LASC Level 1 Specification

Conformance Process for Specification and Acceptance Testing of LASC Communication Protocol

| VERSION: [1.0] | REVISION DATE: [1/10/09] |
| --- | --- |

| Approver Name | Title | Signature | Date |
| --- | --- | --- | --- |

Table of Contents

1.	Introduction	4

1.1	EIP overview	4

1.2	EIP Objects	5

1.3	Session Overview	7

1.4	Message Protocol	7

1.5	NOP Command	9

1.6	ListServices Command	9

1.7	ListIdentity Command	11

1.8	RegisterSession Command	12

1.9	UnRegisterSession Command	13

1.10	SendRRData Command	13

1.11	SendUnitData Command	15

1.12	Full Message Layout	15

2.	EIP server requirements	16

2.1	Mandatory Class Details	17

3.	EIP client requirements	20

3.1	Base Ethernet communications	20

4.	Testing Base Ethernet Communications	20

4.1	Test ICMP	20

4.2	Test UDP request/response	20

4.3	Test TCP request/response	20

4.4	Test 16 concurrent TCP connections	20

4.5	Test  Cable Pull	21

5.	Testing Encapsulation Commands	22

5.1	Test NOP (TCP)	22

5.2	Test List Identity (TCP or UDP)	22

5.3	Test RegisterSession (TCP)	22

5.4	Test UnRegisterSession TCP	23

5.5	Test ListServices (UDP or TCP)	23

5.6	Test SendRRData (TCP)	24

6.	Testing Session Handling/message handling	25

5.7	Multiple sessions	25

5.8	Session Timeout	25

5.9	Message Handling	26

7.	Testing Implementation of common services	28

6.1	Get Attribute Single	28

6.2	Set Attribute Single	28

8.	Testing Implementation of Core Objects	29

5.10	EDS objects checked	29

9.	Testing Error Codes	31

10.	Timeout testing	31

11.	Retest concurrent sessions	31

12.	Appendix A– Data Sizes	32

# Introduction

This document covers the requirements for a device (server) or client to meet LASC Level 1 compliance levels. This covers, in brief, the EIP (Ethernet/IP) stack (Layers 1, 2, 3/4) and a subset of the Application layer functionality (Layer 7). This is the full communication channel between LASC compliant modules.

This layer of certification is to ensure a basic level of information transfer between any server/client pair.

This document does not cover functionality or message responses outside the specific areas noted. That functionality is covered in the specific Level 2 specification for the device.

## EIP overview

LASC communication protocol is built on the foundation of EIP and CIP. Full specifications may be obtained from ODVA; however this document defines the smaller subset of the requirements that must be implemented for LASC compliance.

The EIP protocol is designed for the exchange of time-critical application information. EIP makes use of standard Ethernet and TCP/IP technology to transport CIP (Control and Information Protocol) communications packets.

CIP is a peer to peer object oriented protocol that provides connection between devices and controllers

## EIP Objects

An object consists of the following (See Figure 4-1.1):

- A set of closely related attributes (data)

- A set of behaviours (functionality)

- A set of services (common or object–specific)

- A set of connections.

A Class Attribute is an attribute that is shared by all objects within the same class.

An Instance Attribute is an attribute that is unique to an object instance and not shared by the object class.

Common Services are those whose request/response parameters and required. Behaviours are defined in CIP Common Specification, Appendix A.

A typical object model for a device will look generally like this:

## Session Overview

The normal life of an EIP session would generally involve the following steps:

- A Client opens a TCP connection to <server IP address>, port 44818

- The client can send non-message commands (NOP,ListServices,ListIdentity etc)

- The server MUST respond to commands requiring a response with a message indicating status

- the client sends RegisterSession message

- the server responds with the session id to be used in all future messages

- The client uses SendRRData to encapsulate data/function requests

- the server responds to every message with a response indicating status, data etc

- When the client has finished, it sends an UnregisterSession message

- The client TCP connection can be closed

## Message Protocol

While the CIP specifies both a ‘connected’ and ‘unconnected’ communication, the LASC standard is to use unconnected messages in all cases. These messages are unconditionally point to point.

There are 7 mandatory commands that must be implemented in all devices shown in the following table:

The other optional commands may be implemented. Every device shall accept commands that it does not support and return a ‘command not supported’ status code without breaking the session or TCP connection.

NOTE: Message structure items are coded in little endian unless specified otherwise.

## NOP Command

The NOP command may be used to check TCP connection status. The client or server may generate this message. The receiver MUST not send any response to this message.

Message Structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x0000 | NOP |
| Length (2) | 0x0000 | No data to be attached |
| Session handle (4) | 0x00000000 | Ignored |
| status (4) | 0x00000000 | Must be 0 |
| Sender context (8) | 0x0000000000000000 | Ignored |
| Options (4) | 0x00000000 | Must be 0 |

## ListServices Command

This command is used to determine which encapsulation service classes are supported. There is only one service class defined, so the response message is completely defined.

Message Structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x0004 | List Services |
| Length (2) | 0x0000 | No data to be attached |
| Session handle (4) | 0x00000000 | Ignored |
| status (4) | 0x00000000 | Must be 0 |
| Sender context (8) | 0x0000000000000000 | Any Sender context – usually message number |
| Options (4) | 0x00000000 | Must be 0 |

Response Structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x0004 | List Services |
| Length (2) | 0x001a | Response data portion |
| Session handle (4) | 0x00000000 | Ignored |
| status (4) | 0x00000000 | Must be 0 |
| Sender context (8) | 0x0000000000000000 | Preserved Sender context from request message |
| Options (4) | 0x00000000 | Must be 0 |
| Item Count (2) | 0x0001 | 1 item |
| Item Type code (2) | 0x0100 | Communications service class |
| Item Length (2) | 0x0014 | Length of the item data |
| Version (2) | 0x0001 |  |
| Capability flags (2) | 0x0020 | Supports TCP CIP |
| Name (16) | 434f4d4d554e49434154494f4e530000 | “Communications” |

## ListIdentity Command

This command may be issued by a client broadcast over UDP to determine available EIP servers on the network.

Message Structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x0063 | Identity |
| Length (2) | 0x0000 | No data to be attached |
| Session handle (4) | 0x00000000 | Ignored |
| status (4) | 0x00000000 | Must be 0 |
| Sender context (8) | 0x0000000000000000 | Must be 0 |
| Options (4) | 0x00000000 | Must be 0 |

Response structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x0063 | List Services |
| Length (2) | 0x002d | Response data portion |
| Session handle (4) | 0x00000000 | Ignored |
| status (4) | 0x00000000 | Must be 0 |
| Sender context (8) | 0x0000000000000000 | Preserved Sender context from request message |
| Options (4) | 0x00000000 | Must be 0 |
| Item count | 0x000x | Items to follow |
| Item type code (2) | 0x000c | Must be 0x0c |
| Item Length (2) | 0x00xx |  |
| Protocol version (2) | 0x0000 | Must be 0 |
| Socket Address sin_family (2) | 0x0000 | All following items are defined in the EDS |
| Socket Address sin_port (2) | 0xaf12 | 44818 for EIP |
| Socket Address sin_address (4) | 0x7f000001 | Big-endian 127.0.0.1 |
| Socket Address sin_zero (8) | 0x0000000000000000 |  |
| Vendor id (2) | 0x0000 | 768 for CSIRO |
| Device type (2) | 0x0000 |  |
| Product code (2) | 0x0000 |  |
| Revision (2) | 0x0001 |  |
| Status (2) | 0x0000 |  |
| Serial Number (4) | 0x0000 |  |
| Product name (8) | SPMS | From EDS |
| State (1) | 0x00 |  |

## RegisterSession Command

This message is sent to commence a data transfer session.

Message Structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x0065 | Register session |
| Length (2) | 0x0004 |  |
| Session handle (4) | 0x00000000 | Ignored |
| Status (4) | 0x00000000 | Must be 0 |
| Sender context (8) | 0x0000000000000000 | Any Sender context – usually message number |
| Options (4) | 0x00000000 | Must be 0 |
| Protocol Version | 0x0001 |  |
| Options Flag | 0x0000 | reserved |

Response structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x0065 | List Services |
| Length (2) | 0x0004 |  |
| Session handle (4) | 0x000000xx | Session id to be used for all future messages |
| status (4) | 0x00000000 | Must be 0 |
| Sender context (8) | 0x0000000000000000 | Preserved Sender context from request message |
| Options (4) | 0x00000000 | Must be 0 |
| Protocol Version (2) | 0x0000 | Must be 0 |
| Options Flags (2) | 0x0000 | Must be 0 |

## UnRegisterSession Command

This message is sent to cease a data transfer session. Either an originator or a target may send this command to terminate the session. The receiver shall initiate a close of the underlying TCP/IP connection when it receives this command

Message Structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x0066 | Register session |
| Length (2) | 0x0000 |  |
| Session handle (4) | 0x00000000 | Ignored |
| status (4) | 0x00000000 | Must be 0 |
| Sender context (8) | 0x0000000000000000 | Any Sender context – usually message number |
| Options (4) | 0x00000000 | Must be 0 |

There is to be NO response from this message

## SendRRData Command

Message Structure

| Field | Example (bold is mandatory) | Notes |
| --- | --- | --- |
| CMD(2) | 0x006f |  |
| Length (2) | 0x0018 |  |
| Session handle (4) | 0x00000001 |  |
| status (4) | 0x00000000 |  |
| sender context (8) | 0x0000000000000001 |  |
| Options (4) | 0x00000000 |  |
| Iface Handle (4) | 0x00000000 |  |
| Timeout (2) | 0x0000 |  |
| Item Count (2) | 0x0002 |  |
| Address Type (2) | 0x0000 |  |
| Address Length (2) | 0x0000 |  |
| Data Type ID (2) | 0x00b2 |  |
| Data Length (2) | 0x08 |  |
| Service (1) | 0x0e |  |
| Path Size (1) | 0x03 |  |
| Path (size*2) | 0x206424003003 | Class tag = 0x20 Class=0x64 Instance tag=0x24 Instance = 0 Attribute tag =0x30 Attribute = 3 |
| Data (variable) | none |  |

Response example for the above packet

|  |  | Notes |
| --- | --- | --- |
| CMD(2) | 0x006f |  |
| Length (2) | 0x0016 |  |
| Session handle (4) | 0x00000001 |  |
| status (4) | 0x00000000 |  |
| sender context (8) | 0x0000000000000001 |  |
| Options (4) | 0x00000000 |  |
| Iface Handle (4) | 0x00000000 |  |
| Timeout (2) | 0x0000 |  |
| Item Count (2) | 0x0002 |  |
| Address Type (2) | 0x0000 |  |
| Address Length (2) | 0x0000 |  |
| Data Type ID (2) | 0x00b2 |  |
| Data Length (2) | 0x06 |  |
| Service (1) | 0x8e | Note that the requested service is OR’d with 0x80 in the response |
| Reserved | 0x00 |  |
| Status | 0x00 |  |
| Additional status size | 0x00 |  |
| Data (variable) | 0x7400 | The size of the data returned is DataLength-4-additionalStsSize*2 |

## SendUnitData Command

This command is unused in the LASC subset of the EIP protocol. A Client should not send this message. A server should not reply to this message.

## Full Message Layout

A full message including all the headers is as shown in the image below:

RR Messages have a fixed overhead of 113 bytes + variable size path + variable size data.

Responses have a fixed overhead of 115 bytes + variable size additional status + variable size data.

# EIP server requirements

Level 1 compliance covers the basic communication protocols and required functionality for every device.

Every server implementation MUST have at least the following functionality:

- Base Ethernet communications, including

  - TCP/IP connection, including error handling and connection checking (RFC 793)

  - IP Version 4 (RFC 791)

  - UDP (RFC 768)

  - Address Resolution Protocol (ARP) (RFC 826)

  - Internet Control Messaging Protocol (ICMP) (RFC 792)

  - Internet Group Management Protocol (RFC 1112 + 2236)

  - IEEE 802.3 as defined in RFC 894

- Correct and valid responses to Encapsulation commands

- 0x0, NOP

- 0x4, List Services

- 0x63, List Identity

- 0x65, Register Session

- 0x66, Unregister Session

- 0x6f, Send RR Data

- 0x70, Send Unit Data

- Unconnected message handling Including error handling, message timeouts and status responses

- Implementation of core objects

- Class 0x01 (Identity Object)

  - Class Attribute 1 (Revision)

  - Instance Attributes 1-7

- Class 0xf5 (TCP/IP Interface Object)

  - Class Attribute 1 (Revision)

  - Instance Attributes 1-6

- Class 0xf6 (Ethernet link Object)

  - Class Attribute 1 (Revision)

  - Instance Attributes 1-3

- Implementation of common services

- 0x0e, Get Attribute Single

- 0x10, Set Attribute Single

- Implementation of error responses

- 0x0, success

- 0x1, invalid command

- 0x3, incorrect data

- 0x64, invalid session handle

- 0x65, invalid length

- 0x69, unsupported command revision

- Implementation of message forwarding

- EDS description file for compliance testing.

Other functionality may be described in the Level 2 components of the device specifications.

## Mandatory Class Details

### Identity Object

Vendor shall be 768

Device type shall be 100

The other details shall be vendor-specific and entered in the EDS for the device

### TCP object

The mandatory information in the EDS for this class will be the IP Address, coded in dotted string (e.g. 10.0.0.10) in Attribute 5.

### Ethernet object

# EIP client requirements

Every client implementation MUST have at least the following functionality:

## Base Ethernet communications

TCP/IP connection, including error handling and connection checking (RFC 793)

IP Version 4 (RFC 791)

UDP (RFC 768)

Address Resolution Protocol (ARP) (RFC 826)

Internet Control Messaging Protocol (ICMP) (RFC 792)

Internet Group Management Protocol (RFC 1112 + 2236)

IEEE 802.3 as defined in RFC 894

# Testing Base Ethernet Communications

## Test ICMP

The device is to be issued an ICMP (ping) message. The device should respond with its IP address. The round trip time should be less than 10ms.

## Test UDP request/response

UDP is tested by sending a ListIdentity command (0x0063) via UDP, the device should respond with a valid ListIdentity response. The details for the ListIdentity packet should match the details coded in the EDS file supplied with the device.

## Test TCP request/response

The TCP connection is tested by sending the ListIdentity command (0x0063) via TCP. The device should respond with a valid ListIdentity response, with values populated as per the EDS file provided with the device.

## Test 16 concurrent TCP connections

16 concurrent TCP connections are to be simultaneously established by connecting to port 44818 of the device. Once these connections are established each connection is to have a Register Session command (0x0065) issued via TCP. Each of the sessions will respond with valid connection details. All responses should be identical to that of the tests above. The connections should now be closed.

## Test  Cable Pull

2sec / 120sec resume (for these tests there will be a network switch between the PC and the device being tested)

The user will be prompted by the testing software to remove the Ethernet cable for 2 seconds. The cable is to be disconnected directly from the device under test. After a 2 second pause the cable it to be re-inserted.

The user will be prompted by the testing software to remove the Ethernet cable for 120 seconds. The cable is to be disconnected directly from the device under test. After the 120 second pause the cable is to be re-inserted.

The user will be prompted by the testing software to remove the Ethernet cable for 2 seconds. The cable is to be disconnected directly from the PC. After a 2 second pause the cable it to be re-inserted.

The user will be prompted by the testing software to remove the Ethernet cable for 120 seconds. The cable is to be disconnected directly from the PC. After the 120 second pause the cable is to be re-inserted.

# Testing Encapsulation Commands

## Test NOP (TCP)

The NOP command may be used to test the underlying TCP connection without breaking any existing EIP session. The NOP may be sent by either the client or the server. There is no response to the NOP command.

Prerequisites: No EIP session. Send a NOP message and ensure no response is received.

Prerequisites: EIP session registered. Send a NOP message and ensure no response is received.

## Test List Identity (TCP or UDP)

The ListIdentity message is sent by a client to identify EIP devices available. When sent as a UDP message, it is a broadcast to identify all devices on the subnet. When sent via TCP, it may be used to list the parameters of the device connected on the current TCP connection.

In all the tests below, each of the data items should be checked against the EDS for validity.

Prerequisites: No EIP session. Send a ListIdentity message via UDP. Ensure DUT responds with valid ListIdentity response.

Prerequisites: NO EIP session. Send a ListIdentity message via UDP broadcast. Ensure DUT responds with valid ListIdentity response.

Prerequisites: TCP communications channel. Send a ListIdentity message via TCP. Ensure DUT responds with valid ListIdentity response.

## Test RegisterSession (TCP)

The RegisterSession message is used to commence a session between the client and the server. No messages other than NOP and ListIdentity may be validly used before the session has been started.

Where a response is received, the following items should be validated: Sender context, session handle, status and

Prerequisites: No EIP session. Send a RegisterSession message via UDP. Ensure no response is received

Prerequisites: No EIP session. Send a RegisterSession message via TCP. Ensure a valid response is received.

Prerequisites: No EIP session. Send a RegisterSession message via TCP with Protocol version 2. Ensure a valid response is received with status 0x69 and highest supported version in Protocol version field

## Test UnRegisterSession TCP

The Unregister session is used to terminate the current session. The session is also closed when the TCP connection is lost. There is no response to this message

Prerequisites: No EIP session. Send an UnRegisterSession message. No response

Prerequisites: EIP session. Send an UnRegisterSession message. No response.

Prerequisites: EIP session. Send an UnRegisterSession message. Send another message with same session if. Error response

Prerequisites: No EIP session. Send an UnRegisterSession message with wrong session id. No response. Send other message with correct session id. Should still respond ok.

## Test ListServices (UDP or TCP)

The ListServices message is used to identify which encapsulation service classes are supported (not the Explicit Messaging services such as Get_Attribute_Single) Currently, only one service is defined by EIP and it indicates that the device supports encapsulation of CIP packets.

Prerequisites: No EIP session. Send a ListServices message via UDP. Valid reply message from DUT.

Prerequisites: No EIP session. Send a ListServices message via TCP. Valid reply message from DUT.

Prerequisites: EIP session. Send a ListServices message via UDP. Valid reply message from DUT.

Prerequisites: EIP session. Send a ListServices message via TCP. Valid reply message from DUT.

## Test SendRRData (TCP)

This message is the main transfer of encapsulated request/reply packets between the client and server. This message, for Landmark compliance purposes, is used solely to transfer UCMM (unconnected) messages as per the format described in section 1.10.

Details of required data for UCMM testing is in section 9.

Prerequisites: No EIP session. Send a SendRRData message via TCP. With class: 1, instance: 1, attribute: 1, service: 14. No reply message from DUT.

Prerequisites: EIP session. Send a SendRRData message via TCP. With class: 1, instance: 1, attribute: 1, service: 14. Valid reply message from DUT.

# Testing Session Handling/message handling

This section describes the testing required to ensure that the device or application under test correctly handles messages and sessions.

## Multiple sessions

Test 16 registered sessions simultaneously – each session should be given a unique session id.

Test RRData in each connected session – each message should receive a valid response, including correct session id and message context

## Session Timeout

Test connected session timeout with no messages – the session should not have disconnected after 2 seconds.

Test connected session timeout with no messages – the session should have disconnected after 120 seconds.

## Message Handling

Send truncated message, test message timeout – no response received.

Send Long messages (dual message), check response received for both messages – should receive valid responses.

Send garbled message (shorter than normal packet) – should get an error return

Send garbled message (same size as normal packet) – should get an error return

Send garbled message (longer than normal packet) – should get an error return

Send 1000 EIP messages with no wait time (stress test) –should get valid responses for each message

send messages with wrong session identifier – should get error message.

send messages with negative/same sender context – should get sender context copied and returns no error

send messages with previous (and closed) session handle -should get error message

send register, TCP fail then reconnect (cable pull pc side), messages should fail

send register, TCP fail then reconnect (cable pull pc side), unregister message should fail

send register, TCP fail then reconnect (cable pull device side), messages should fail

send register, TCP fail then reconnect (cable pull device side), unregister message should fail

send messages: nop, register, nop, listservices,nop, unregister – test nop does not interfere with normal operation

# Testing Implementation of common services

## Get Attribute Single

Class 1, instance 1, attribute 1 is read using SendMessage

## Set Attribute Single

This will involve the write of an attribute.

# Testing Implementation of Core Objects

## EDS objects checked

Each Attribute in the EDS file for the device is checked for a valid response from the device.

If the attribute is implemented, then the General Status response should be 0x00 (Success).  If the attribute is not implemented, then the General Status Response should be 0x05 (PATH destination unknown).  If the attribute is not part of the object definition then the General Status Response should be 0x14 (Undefined attribute).  If the proper General Status Response of 0x00 (Success) is returned then the size of the attribute is verified as well as its value.

The EDS file should consist of at least the following.

### Identity Object

#### Test Attribute 1: Vendor Id

#### Test Attribute 2: Device Type

#### Test Attribute 3: Product Code

#### Test Attribute 4: Revision

#### Test Attribute 5: Status

#### Test Attribute 6: Serial Number

#### Test Attribute 7: Product name

### TCP/IP Object (Class Code 0xf5)

#### Test Attribute 1: Revision (implemented)

For instance 1:

#### Test Attribute 1: Status

#### Test Attribute 2: Configuration Capability

#### Test Attribute 3: Configuration control

#### Test Attribute 4: Physical Link Object

#### Test Attribute 5: Interface Configuration

#### Test Attribute 6: Host Name

#### Test Attribute 7-99 (optional)

For instance 2:

#### Test Attribute 1, response should be 0x5 (path destination unknown)

### Ethernet Link (Class Code 0xf6)

#### Test Attribute 1: Revision (implemented)

For instance 1:

#### Test Attribute 1: Interface Speed

#### Test Attribute 2: Interface Flags

#### Test Attribute 3: Physical address

#### Test Attribute 4-99 (optional)

# Testing Error Codes

The correct error code should be provided in the status code field in the following situations

Error 0x1: invalid or unsupported encapsulation command

Error 0x3: Poorly formed or incomplete data in the data portion

Error 0x5: Path destination unknown

Error 0x64: invalid session handle

Error 0x65: invalid length message

Error 0x69: unsupported encapsulation protocol revision

# Timeout testing

List Identity Response/UDP < 250ms

List Services Response/TCP (assuming existing TCP connection)< 250ms

Unconnected explicit response - TCP connection established (Specific internal object/attribute would be tested)< 100ms

Two back-to-back explicit message requests without dropping either > 1ms

# Retest concurrent sessions

16 concurrent TCP connections are to be simultaneously established by connecting to port 44818 of the device. Once these connections are established each connection is to have a Register Session command (0x0065) issued via TCP. Each of the sessions will respond with valid connection details. All responses should be identical to that of the tests above. The connections should now be closed.

# Appendix A– Data Sizes

| Version | Description of Change | Author | Date |
| --- | --- | --- | --- |
| 1.0 | Created | Mark Dunn | 4/2/08 |